F529U

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Board ID | Doc Version | Date | Author | ASIC ID | Production |
| F529U | 0.1 | 10.05.2022 | L. Ghislotti | 529 | US |

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | **#** | **Test Description** | **Pass** | **Note** | |
| Visual | 1 | Visual inspection | Yes | See notes (if any) | |
| Bias | 2 | Analog | Yes | AVDD [V] | 1.797 |
| Yes | IVDD [mA] | 134.2 |
| Digital | Yes | DVDD [V] | 14.190 |
| Yes | IDVDD [mA] | 5 |
| Calibration | Yes | 3V3 [V] | 18.370 |
| Yes | I3V3 [mA] | 3 |
| Bias current | Yes | Ibias [mA] | 4.92 |
| Reference Voltages | Yes | VCMSH [V] | 1.199 |
| Yes | VCM [V] | 0.884 |
| Yes | RVCM [V] | 0.899 |
| Digital | 3 | Configuration test | Yes |  | |
| Wrong address test | Yes |  | |
| Temperature Sensor | Yes | ADC | 1085 |
| T [°C] | 29.584 |
| Analog | 4 | Input-output characteristic | Yes | # of non resp. ch | 0 |
| 5 | Noise ENC FWHM @ tp=6 | Yes | Ch 0 [keV] | 6.467 |
| Ch 7 [keV] | 26.471 |
| Ch 15 [keV] | 7.277 |
| Ch 16 [keV] | 17.202 |
| Ch 23 [keV] | 27.355 |
| Ch 31 [keV] | 12.229 |
| 6 | Threshold dispersion @ tp=6 | Yes | Before FT [DAC] | 3.636 |
| After FT [DAC] | 1.409 |
| 7 | Pedestal dispersion @ tp=6 | Yes | Dispersion [ADC] | 8.115 |

**Notes:**

scratches on top  
scratches on bottom